BSW-M & BSW-F & BSW-P

Calibration and setting

1. Parameter setting

1-1 Press power switch to start the scale. While it display 1-9, Press the key  ,,  one by one. When it shows ‘F1-diu’,it means enter into the setting.

1-2 Press the key  or  to select options and parameter.

Press the key  to enter selected option or confirm setting.

| Display | Option | Parameter |
| --- | --- | --- |
| F1-diu | Division | 1 2 5 10 20 50 |
| F2-Fnt | Decimal point | - 0.0 0.00 0.000 0.0000 0.00000 |
| F3-CAP | Capacity | 3000～6000 |
| F4-CAL | Calibration |  |
| F5-tra | Zero-tracking | - 0.5 1.0 1.5 2.0 2.5 3.0 3.5 4.0 |
| F6-FLt | Filtration | 0～16 |
| F7-Lit | Backlight switch | Au on off |
| F8-bud | Baud rate | 2400 4800 9600 19200 |
| PRINT MODE |  | 9600 |
| HI-LOW limit mode |  | 2400 |
| F9-ode | Print mode | 0 1(Hand) 2(Auto) |
| PRINT MODE |  | 1 or 2 |
| HI-LOW limit mode |  | 0 |
| Fa-off | Time for automatic off | 5 10 15 20 25 30 dis |
| Fb-ALr | Alarm for hi-low limit | in out off |
| PRINT MODE |  | off |
| HI-LOW limit mode |  | In or out |
| Fc-dny | Movement lock | 5 10 15 20 25 30 35 40 45 |
| Fd-unt | Unit switch | U1 ON U2 ON U3 ON Def U1 U2 U3 |

2.Calibration:

When it shows ‘F4-CAL’, press the key  .It will shows ‘CAL-1’,which means enter into calibration setting.

2-1 Press the key  again to show ‘0.0’,then input the value of calibration weight.

2-2 Press the key  again. When it shows ‘Put’, place the calibration weight on the pan. Press the key  to confirm. It will show ‘F4-CAL’，which means finish setting.

2-3 Press the key  to exit from the calibration mode.