UCA-M Calibration and setting

1. Parameter setting

1-1 Press the key  to start the scale. When it count from 111 to 999, promptly input password ‘759’. When three windows show  ‘diu   XXXX

F01’,

it means enter into the setting.

1-2 Press the key  or  to select options and parameter.

Press the key  to enter selected option or confirm setting.

| Display | Option | Parameter |
| --- | --- | --- |
| diu-F01 | Division | 1 2 5 10 20 50 |
| Pnt-F02 | Decimal point | - 0.0 0.00 0.000 0.0000 |
| CAP-F03 | Capacity | 3000～6000 |
| CAL-F04 | Calibration |  |
| tra-F05 | Zero-tracking | - 0.5 1.0 1.5 2.0 2.5 3.0 3.5 4.0 |
| FLt-F06 | Filtration | 0～9 |
| Lit-F07 | Backlight switch | Au on off |
| Bud-F08 | Baud rate | 1200 2400 4800 9600* 19200 （9600,for connect to printer and PC） |
| Ode-F09 | Print mode | 0 not print 1 press “P” to print 2 auto print when weighing 3 continuous print |
| Off-F10 | Time for automatic off | 5 10 15 20 25 30 dis |

2.Calibration:

When it shows ‘CAL-F04’, press the key  .It will shows ‘CAL-1’,which means enter into calibration setting.

2-1 Press the key  again to show ‘0.0’, then input the value of calibration weight.

2-2 Press the key  again. When it shows ‘Put’, place the calibration weight on the pan. Press the key  to confirm. It will show ‘F4-CAL’，which means finish setting.

2-3 Press the key  to exit from the calibration mode.

3. RS232 connection

3-1 Specifications:

Baud rate : 9600

Parity   :none

Data bit  :8

Stop bit  :1

3-2 Print form:

3-3 RS232 connection

PORT OF SCALE                   PORT OF COMPUTER (PRINTER)

TXD-----------------------------------------------RXD

RXD-----------------------------------------------TXD

GND-----------------------------------------------GND