Electrical Balance/BB Series

Operation Manual

Specification

| Capacity | 100g | 200g | 300g | 600g | 1200g |
| --- | --- | --- | --- | --- | --- |
| Division | 0.01g | 0.01g | 0.01g | 0.02g | 0.05g |
| Pan Size | ￠123mm |  |  |  |  |
| Power | AC voltage:110v/220v; DC voltage 4.5v(3 500mA cell) |  |  |  |  |

Chapter 1 Functions and Specifications

1-1  Main function:

*Unit switching :gram, kilogram；

*Single screen simple counting function；

*Adopting high precision sensors,high precision 1/30000, accurate stable；

*LCD back-light and easy to read

*stainless steel scale clean easily

*RS232 interface which can connect to computer,printer etc equipment

1-2 Specification

| Max. Capacity | 100g | 200g | 300g | 600g | 1200g |
| --- | --- | --- | --- | --- | --- |
| Division | 0.01g | 0.01g | 0.02g | 0.02g | 0.05g |
| Mini.Capacity | 1g | 1g | 1g | 2g | 4g |
| Accuracy class | Ⅲ |  |  |  |  |
| pan size | ￠123mm |  |  |  |  |
| Power | AC voltage:110v/220v; DC voltage 4.5v( 3 500mA cell)P.S.AC/AD can not be used at the same time |  |  |  |  |

Chapter  2   Installation

- Stainless steel pan

- Screw M4*20 SB

- plastic pan (round)

- Use wrench to lock the screw after the plastic pan put on.

2.Then put on the stainless steel pan

Chapter  3  Display and keyboard

3-1 display

    - →0←(zero indication)–when ▼appears on the above position, said zero.

    - NET (net weight)- when ▼appears on the above position,said container

    - gross weight has been deducted .

    - (counting)- when ▼appears on the above position,said the quantity or percentage.

    - （stable）- when ▼appears on the above position,said weight in a stable state.

3-2  keyboard functions

:print function (available)

:for selecting various unit of weight

:return to zero and remove the peel key

:the parts unit and percentage key

:Choose the format of printing(used when connected to printers)

Under the state of weight,keep pressinguntil display no 01, then press    to choose the format of printing,then press to sure.

Chapter 4 operating instructions

4-1  Getting start

4-1-1 select a suitable work ares, make sure that the scale feet should level and under the state of no-wind.

4-1-2 turn on the power ,the scale will return to zero ,then finish the start.

4-1-3 make sure it is full of power or not.

4-1-4 make sure the cursor ▼point to or press the key    to return to zero

4-2    counting function

4-2-1 take samples first(10,20,50,100 or 200,the more the better) on the pan,press ,the display will show 10,20,50,100,200 by cycle,press when it shows the quantity you put,the display will the quantity you put.Then put samples you want on the pan,it will show the quantity.

4-2-2 if want to know total weight ,press,again press return to counting state .

4-3   Calibration(this can be used after the operation FA-CAL)

Chapter 5  BB balance calibration and setting

When power turn on,self test.press three times，enter setting, means select parameter,means confirm setting

.

| Display | Option | Parameter |
| --- | --- | --- |
| F1-diu | Division | 1 2 5 10 20 50 |
| F2-Fnt | Decimal point | - 0.0 0.00 0.000 0.0000 0.00000 |
| F3-CAP | Capacity | For example:300g,300.00 |
| F4-CAL | Calibration | Please see attached 1 |
| F5-tra | Zero-tracking | - 0.5 1.0 1.5 2.0 2.5 3.0 3.5 4.0 |
| F6-USr | Calibration | Please see 4.3 |
| F7-Lit | Back-light set | Au on off |
| F8-bud | Baud rate | 2400 4800 9600* 19200 |
| F9-ode | Print mode | 1.manual 2.auto 3.continuous 4. ULP .manual 5. ULP auto |
| FA-unt | Unit switch Open default unit set | U1(g) U2(kg) U3(1b) Def U1 U2 U3f |
| Fb-ALr | high/low limit alarm set | IN OUT OFF |

Attached 1

When show “F4-CAL”,

Press it will show CAL-1(if it show CAL-3 means 3 divisions calibration,can press to select)

Press ,it will show 005.00,press to select parameter

Pressto sure,when show “put”,put on the calibration parameter,pressto sure,then finish the calibration.

Chapter 6 Recharge, Maintenance &Repair

6-1  Maintenance

6-1-1  Please take good care of the BB scale for it is a precision instrument

6-1-2  Place an object on the pan softly in less to destroy load cell for the power of moment hit is too heavy.

6-1-3  In order to keep the scale in formal, please recharge the battery in time

6-2 Repair

6-2-1  Please don’t open the scale without permission.

6-2-2  Please contact with maintenance-station directly when it doesn’t work

Chapter 7  Common fault and debug method

| Common fault | solution |
| --- | --- |
| Can’t boot up | 1.battery no power 2.cell line virtual welding 3.display and AD board connection off |
| Can’t charge | 1.charging line virtual welding 2.diode virtual welding(D2-D7),resistance virtual welding (R8,R9) 3.triode broken (7808) 4.battery can’ t storage power |
| Keyboard problems | 1.keyboard broken 2.K\B stand bad contact or virtual welding 3.I\C broken |
| Display not clear and hyphenation | 1.LCD problems .short circuit 2.display chip short circuit 3.7150 triode broken |
| Weighing inaccurate | 1.electrostatic,wind and other environmental factors 2.scale on stable, four feet off balance 3.the stand slanted lead to load cell run into the shell |

Chapter 8  spare parts list

| Name | Quantity |
| --- | --- |
| Main board | 1 |
| Display board | 1 |
| Charger | 1 |
| pan | 1 |
| Frame | 1 |
| Load cell | 1 |
| Plastic shell | 1 |
| instruction | 1 |