UT-C

Calibration and setting

1. Parameter setting

1-1 Press power switch to start the scale. While it display 1-9, Press the key  ,  ,  one by one. When it shows ‘F1-diu’,it means enter into the setting.

1-2 Press the key  or “PEAK” to select options and parameter.

Press the key  to enter selected option or confirm setting.

| Display | Option | Parameter |
| --- | --- | --- |
| 1-diu | Division | 1 2 5 10 20 50 |
| 2-Fnt | Decimal point | - 0.0 0.00 0.000 0.0000 0.00000 |
| 3-CAP | Capacity | 3000～6000 |
| 4-CAL | Calibration |  |
| 5-tra | Zero-tracking | - 0.5 1.0 1.5 2.0 2.5 3.0 3.5 4.0 |
| 6-FLt | Filtration | 0～16 |
| 7-Lit | Backlight switch | 0 1 2 |
| 8-bud | Baud rate | 2400 4800 9600 19200 |

2.Calibration:

When it shows ‘F4-CAL’, press the key “PEAK” .It will shows ‘F4-CAL’,which means enter into calibration setting.

2-1 Press the key  to show ‘00000’,then use the key and “PEAK” to input the value of calibration weight.

2-2 Press the key  again. When it shows ‘Put 1x’, hang on the calibration weight. Press the key  to confirm. It will show ‘F4-CAL’，which means finish setting.